Double SPA waveform Viewer sample software 　　　　　V 1.0

Advantest corporation

- Install double SPA waveform Viewer sample software on PC.

Run Setup.exe of InstDoubleSPA, sample software.

- Set up GPIB address of U3751 series Spectrum analyzer.

SYSTEM –Remote Control –GPIB Address from panel key.

The factory default is 8. Modify the address to the appropriate values for your configuration.

（example: to change the address to 10, press １　０　Enter (Hz key)

- Set GPIB address for the second SPA.

The second Spectrum Analyzer should have a different address than the First.

For example, if the first one analyzer is set to 8 (the default) then set the GPIB address of the second to anything valid address except 8.

Then connect the PC, the first and the second Spectrum Analyzer with GPIB cable.

- Run the sample software.

START – PROGRAM - U3700 Screen Viewer - U3700 Screen Viewer

Set the GPIB address of the Spectrum Analyzers. (In this example, 8 and 10)

Check the “SPA 1,2”, and the waveforms from both Spectrum Analyzers will be displayed.

The radio buttons determine the Active Spectrum Analyzer in the waveform view. Both waveforms will be visible but the changes made to the parameters will be applied to the active instrument

- Press Connect

When the connection between the PC and Spectrum Analyzer been established, “Connected” will be displayed on message line located at the bottom part of window. If you have checked Monitor, the communication log will be displayed.

The parameters of the Spectrum analyzer can be controlled manually using the “Sending Text” command line. Type a GP-IB command in “Sending Text” and push “Send Text” key. The U3751 should respond accordingly. Send a query command (e.g. CF?) and you can acquire data from the instrument that will appear in the monitor window.

If you press SendText with the above mode（＊IDN）, the information on the Spectrum Analyzer’s identity will be displayed.

Note: When you switch between Spectrum analyzers a reconnection is required.

Before executing Waveform View, uncheck the Monitor check box.

- Waveform Viewer

Press “Waveform View”

You will find that the active configurable Spectrum Analyzer listed in “Active SPA:” In the example SPA-1 is shown as the active spectrum analyzer. If you change parameters on this display they will be applied to the active spectrum analyzer.

The waveform trace of the Active spectrum analyzer (in this example SPA-1) is displayed in yellow and the second spectrum analyzer trace (SPA-2) is in green. The date and time displayed at the bottom right is the system time from the PC.

Most of the key parameters can be set in this screen. Other parameters not available in this screen should be set by returning to the start up screen and entered using the command line.

The test parameters are initially set to 0. After the parameters have been entered, execute a measurement and the parameters will be sent to the active spectrum analyzer. To execute a measurement, press “SWP Cont” or “SWP Single”.

In the lower right are additional controls.

“BMP” creates a bit map file of this screen in the PC.

“Preset”  resets the active spectrum analyzer to factory settings.

“CLOSE” closes the Waveform Viewer and returns to the start up page.

- Exiting program

Exit the Waveform Viewer by choosing “CLOSE”.

Terminate the GPIB connection by choosing “CLOSE” in the start up screen, and exit the program by choosing “EXIT”.

Additional information:

Screen display using 1 unit (SPA1)

When there is one spectrum analyzer, make sure the “SPA 1,2” box is not checked, select and connect with Active spectrum analyzer, and then press Waveform View.

Displaying 2 spectrum analyzers in separate screens

Launch the program again so that you have 2 copies running and setup the spectrum analyzer as described above in “Screen display using 1 unit (SPA1)”. This time select the SPA-2 radio button as the Active spectrum analyzer, and then press Waveform View.

Please insure that the GPIB addresses are not the same.

Note: Side-by-side display